Tutorial AQUAStalk software
Measuring protocol :
❶ Fill battery of Aquascat and make sure it works properly by connecting the computer
❷ Temperature
❸ Depth of transducer
❹ Collect profile of both fine and sand sediments sample for the calibration, where sand sample concentration (close to the bottom) is maximum.
❺ Collect profile of both fine and sand sediments sample for the validation, at a different position compared to the calibration profile.

---
Tutorial AQUAStalk software
Step 0 : Click on Connect button to start and connect the logger.
If Aquascat is ok then the connection of logger is performed automatically.

---
Settings
Step 1 : Click on Deploy
Step 2 : Select frequency for transducer(s)
3
2
Channel 1 = mark with a red dot
Other channel numbered clockwise
1
4

---
Step 3 : Project description
(it’s not the project name)
Step 4 : Define the measurements time
Click on Correct Logger Time to set the measurements time to PC time (=UTC time)

---
Step 5 : Select the measurements time
Step 6 : Define gains

---
Step 7 : Define gains
Step 8 : Define Cell size
Adjust the cell size to the water depth by checking the sentence in « Stored Profile » :
   Profile start at 2,00cm end at 514,00cm
Always select +20dB

---
Step 9 : Define Pulse length
Step 10 : Define Transmit Power
Keep the same value as « Bin size »
To evaluate Bin size in time (sec) :
Tx Pulse Length x Frequency
Select 0dB for emission + reception (measurements use to compute sediments concentrations)
Select OFF for reception only (measurements of noise)

---
Step 9 : Define number of bins
Step 10 : Define Transmit Power
Keep the Start Bin at 1 and Number of Bins at 256 (max)

---
Step 11 : Define measurement mode
Step 12 : Auto Logging
Auto Logging : Aquascat is triggered at a set time
Trigger Only : Aquascat is trigger by the user
Set the time when the measurement are triggered

---
Step 11 : Trigger Only
Step 12 : Auto Logging
Set the time when the measurement are triggered (internal trigger).
Set waiting time between two recordings.
The measurements are triggered by the user (external trigger)

---
Step 13 : Select Profile Framing Mode to Auto
Step 14 : Choose measurements frequency

---
Step 15 : Number of profiles use to average each profiles (keep 1)
Step 16 : Check the length of measurement in Burst length
Uncheck box

---
Step 17 : Write setting in the actuel measurement project (logger)
Step 18 : Project information
Click on button « Real Time Display » to launch the measurements

---
Measurements
Step 19 : Real Time View
Click on « Start Logging » to start measurements
Modify Settings
Clear plots of Real Time View
Click on « Stop Logging » to stop measurements
Go to the tab « ABS scrolling »

---
Measurements
Step 19 : Real Time View
Step 19 : Real Time View
Measurementsfilename
Measurements timer
(see burst length - step 16)
Click on « Log To File » to record settings and measurements to be taken

---
Measurements
Step 19 : Real Time View
Step 19 : Real Time View
River bed can’t be displayed because maximum depth of the stored profile (see step 8) is not large enough.
Real Time View displays a maximum of 999 profiles.

---
Measurements
Step 19 : Real Time View
Interferences between AQUAscat and another acoustic tool

---
Measurements
Step 20 : Save data
Step 20 : Save data (List of folder)
Click on « Upload Data » to save measurement to file

---
Measurements
Step 20 : Save data (Number of files in selected folder)
Step 20 : Save data (List of files)
Click on « Upload Selected Files » to open a dialog box to save measurements

---
Measurements
Step 20 : Save data (Dialog box)
Step 20 : Save data
Do not enter a filename to keep default filename : YYYYMMDDHHMMSS.aqa

---
AcouSed to check measurements

---
AcouSed to check measurements
❶
ABS (Acoustic Backscatter System) system choice: AQUAscat or UBSediFlow
❷
Downloaded acoustic file: AQUAscat (.aqa) or UBSediFlow (.udt)